# Comparative Analysis of Spline Based Methods for Breathing Effort Estimation in Pressure Support Ventilation

Abstract

Model-based methods of estimating patient breathing effort aim to provide continuous non-invasive estimates of both patient respiratory effort and respiratory mechanics without disrupting tidal breathing. Assumptions about the shape of the breathing effort profile are incorporated into existing models with the intent to reduce the set of possible solutions to the model without excluding the desired physiological solution. A novel spontaneous breathing estimation method has been developed – coined the B-Spline Mixed Integer Quadratic Programming (B-MIQP) method – extending upon existing models to optimize model accuracy and clinical relevance. The B-MIQP method and existing B-Spline method are validated against gold standard esophageal pressure-based measures considering two separate patient cohorts receiving Pressure Support Ventilation (PSV). For completeness, additional comparative analysis of existing Constrained Optimization and Sparse Optimization methods is also considered. Root mean square error of model fits between B-MIQP (0.4 [0.3, 0.6]) and B-Spline (0.8 [0.5, 1.2]) demonstrate no loss of model fitting accuracy with the implementation of additional constraints. Inspiratory respiratory mechanics are within physiological bounds, and model outcomes align ventilator unloading expectations. Statistically significant relationships (p-value <0.05) between esophageal pressure identified and model derived Pressure-Time Product values were observed for the majority of patients, considering both models. The B-MIQP and B-Spline methods may be able to track breath-to-breath changes in patient effort, where performance is shown to depend on the information content of the input signals. A shared model limitation related to the control feedback structure of pressure support ventilation was identified. Additional analysis of other existing model-based methods provided evidence this identified limitation may be shared across most existing methods.

Keywords: Mechanical ventilation, Respiratory monitoring, PSV, Patient effort, Respiratory modeling

1. Introduction

Lung and Diaphragm Protective Ventilation (LDPV) is of growing clinical interest when providing mechanical ventilation care [1]. Titration of spontaneous patient breathing (SB) efforts to avoid vigorous breathing efforts (Patient-Self Inflicted Lung Injury) [2] while maintaining sufficient activation of the diaphragm [3, 4] is imperative to achieve LDPV. Considering Pressure Support Ventilation (PSV), the ventilator support level can be adjusted by the clinician to adequately unload the respiratory muscles, targeting efforts consistent with quiet resting breathing [5]. Measures of patient breathing effort inform the clinician of patient state when performing these adjustments. However, current measures of patient effort are not easily accessible, and, as such, ventilator settings may be inappropriate for the patient, leading to injury.

Esophageal pressure (Pes) based measures, such as the Pressure-Time-Product, PTP, are currently considered the gold standard for measuring patient effort [6] and extensively used as a reference measure when validating other estimates of patient effort [7, 8, 9]. However, clinical use is limited, due to concerns about the accuracy of the measures [10], as well as a non-trivial additional clinical burden to establish, monitor, and interpret the measurements [11]. Alternative non-invasive measures either lack accuracy [12], disrupt tidal breathing, [8, 13, 14] and/or cannot be separated from the work of the ventilator [7, 15]. Hence, there exists a need for a non-invasive, continuous, and accurate measure of patient effort that does not disrupt tidal breathing.

Model-based methods have previously been developed to address this need [16, 17, 18, 19, 20]. These methods estimate spontaneous breathing pressure profiles and respiratory mechanics via minimization of a cost function. Fundamentally, this function describes the difference between the measured airway pressure data and a modeled airway pressure signal [21], using measured flow and volume signals as model inputs. As discussed in [19], the airway pressure model is not identifiable, so multiple non-physiological solutions exist. To ensure identifiability, existing model-based methods make assumptions about the shape of the patient’s breathing profile, reducing the solution space and removing non-physiological solutions. Of note, [22] presents a model where the SB effort profile is described as a linear combination of equally spaced basis spline (b-spline) functions, scaled by a corresponding set of spline coefficients. In addition, [16] presents a Sparse Optimization (SO) method that assumes a smooth SB profile, penalizing solutions where the trajectory of the profile changes suddenly. More specifically, a penalty function is considered that penalizes modeled breathing efforts where the second derivative of the signal is non-sparse. By comparison, the Constrained Optimization (CO) [19] and Mixed-Integer Quadratic Programming (MIQP) [20] approaches impose stricter assumptions about the SB profile. These methods assume the SB profile should be monotonically increasing, monotonically decreasing and subsequently constant in three defined regions. The boundaries of these regions are determined respectively via a grid search method and with the use of binary variables that are solved for within a mixed integer optimization framework.

A novel method – referred to as the B-MIQP model – that combines the existing B-Spline and MIQP models, with the aim to improve estimation accuracy, was recently presented in [23]. This method fitted a model over a number of breaths, similar to the B-Spline model. This allowed for variations in model input signals from breath-to-breath to be captured as changes in SB effort rather than changes in respiratory mechanics. In addition, to reduce the computational burden when describing SB effort over multiple breaths, a set of b-spline functions were used to uniquely describe each effort. The corresponding scaling coefficients of the b-spline functions were constrained within a set of monotonic regions. The boundaries of these regions were identified within a mixed integer framework, allowing for identified SB efforts to continue into the ventilator expiration phase. Preliminary testing on a limited patient cohort suggested that the model could track changes in patient effort when evaluated against the gold standard, but also highlighted some limitations related to PSV that may hinder estimation accuracy.

In this study, the B-MIQP model has been validated on a larger patient data set against the gold standard measure. The performance of the model is compared to that of the B-Spline model. For completeness, a brief comparative analysis with the Constrained Optimization and Sparse Optimization approaches is also considered. Comparison of model outcomes suggest that accuracy of model estimates is linked to the degree of variation in the airway pressure signal. Further, we reconfirm that the mechanisms of pressure support ventilation may partially limit model identification by obscuring some pulmonary mechanics, a limitation that was observed to impact all considered identification methods.

2. Methods

2.1. Clinical Data

Data was collected from two separate patient cohorts. First, three (3) patients (male, aged 57-83 years) were recruited at the Civil Hospital of Brescia, Italy (referred to as Cohort 1). This cohort represents a subset of a larger prospective cohort, currently being recruited as part of a separate trial on LDPV. Preliminary analysis of the B-MIQP model was conducted considering Cohort 1 in [23]. Airway flow, pressure, and volume measurements were recorded by the Servo-U ventilator (Maquet Critical Care, Solns, Sweden), sampled at 100 Hz. Esophageal pressure was obtained via a NutriVent nasogastric catheter (Sidam, Mirandola, Italy).

Second, an additional six (6) spontaneously breathing patients (4 male, 2 female, aged 52-68 years) were recruited at the University Hospital of Heraklion Intensive Care Unit, Greece (referred to as Cohort 2). Patient data was collected as part of routine clinical monitoring and has been previously considered for analysis of parametric patient breathing models [24]. Airway flow measurements were taken at the patient’s airway opening using a heated pneumotachograph (Hans-Rudolf 3700, Kansas, USA). Airway pressure was measured between the endotracheal tube and the patient circuit Y-piece using a differential pressure transducer (Micro-Switch 140 PC, Honeywell, Ontario). Esophageal pressure measurements were obtained via separate differential pressure transducers (Micro-Switch 140 PC, Honeywell, Ontario) connected to thin latex

(a) Signal filtered using filter adapted from [29].

(b) Signal smoothed using moving average filter.

Figure 1: Esophageal pressure (left) and frequency spectra (right) before and after filtering approaches applied. Pleural pressure (fPpl) and cardiogenic oscillation (fCGO) fundamental/harmonic frequencies identified by black arrows

balloon-tipped catheters. All measurements were sampled at 200 Hz (Windaq Instruments, Ohio, USA)

For both patient cohorts, patients were ventilated under Pressure Support Ventilation (PSV) following local clinical guidelines. The obtained datasets are de-identified. Both studies were purely observational, with no change to care.

A total of 20 breaths per patient were examined to evaluate and compare models. Relevant ventilator settings during periods where breaths were analyzed are listed in Table 1. For Patients 6 and 9, 19 and 18 breaths were considered, respectively, due to the scarcity of suitable breaths for evaluation. For Patients 1 and 2, two sets of 20 breaths were considered at two different support levels to take advantage of a greater data availability.

For both cohorts, end expiratory occlusions were periodically executed to ensure the swing in airway pressure (△Paw) is accurately reflected by the swing in esophageal pressure (△Pes), ensuring correct establishment of the esophageal pressure catheter. Specifically, the △Pes/△Paw ratio during end-expiratory occlusions was maintained between 0.8-1.2 to minimize △Pes inaccuracies. Breaths for analysis were selected either immediately prior or post an end-expiratory occlusion where this ratio was recorded. To further minimize erroneous results, data was excluded where sudden positive spikes in the Pes signal occurred, severe patient-ventilator asynchronies occurred frequently, and/or where the end-expiratory volume was larger than 50 [mL]. Finally, for each breath, the first 50 [ms] of ventilator expiration data was removed to mitigate the impacts of the unmodeled cycling-off ventilator mechanics, consistent with the method procedure outlined in [22].

2.2. Problem Statement

The pressure at the patient’s airways (Paw) is modeled by the widely used single compartment lung model [21], where SB patient effort contributions to overall Paw are described by an appended singular pressure source (Pp):

Pawt=EVt+RVt+PEEP+Ppt,#1

Figure 2: Example B-MIQP and B-Spline fittings considering Patient 2a and 7.

where E and R are patient-specific respiratory system elastance [cmH​2O/L] and resistance [cmH​2O/L/s] constants governing the pressures generated by the tidal volume (V) [L] and flow (V) [L/s] signals, respectively. Positive end-expiratory pressure (PEEP) is the elevated pressure at the end of expiration [25].

As addressed in [19], the solution set for [E,R,Pp] is not unique. Two physiologically meaningful solutions exist, as well as multiple non-physiological solutions. Considering the two physiological solutions, the model can first be used to identify the pressure generated in the pleural space to inflate the lungs (Ppl):

Pawt=ELVt+RVt+PEEP+Pplt,#2

where EL is the elastance of the lungs.

Second, the contributions of the chest wall can be considered to identify the total pressure generated by the patient’s respiratory muscles during a breathing effort (Pmus):

Pawt=ETVt+RVt+PEEP+Pmust,#3

where ET describes the total elastance of the patient’s respiratory system, considering both chest wall elastance (ECW) and lung elastance (EL):

ET=ECW+EL.#4

In general, by rearranging Equation (1) to solve for Pp, any set of values for [E,R] can be chosen and a corresponding Pp can be identified to yield a perfect model fit. As such, without P​p shape assumptions incorporated into a model, multiple non-physiological solutions will exist, leading to poor practical identifiability.

The model constraints imposed by the B-Spline and B-MIQP models, as introduced in the following Section 2.3, allow for convergence to either physiological solution under different experimental conditions. In this study, model outputs were compared to a surrogate measure of pleural pressure. In certain scenarios, the model solution may instead reflect the muscle pressure solution in Equation (3) rather than the pleural pressure solution in Equation (2). However, given Ppl represents a component of the larger Pmus, findings of comparisons of model outputs to Pes can be extrapolated to changes in Pmus. The experimental conditions that lead to different physiological solutions are considered in Section 4.

2.3. Models

2.3.1. B-Spline Model

To regularize the problem, the B-Spline model reduces the number of parameters used to describe SB efforts. Rather than describing efforts at each discrete point in time, the SB effort profile (Pp) is described as a linear combination of b-splines [22]. M=Tmax/kw+d number of 2nd order b-splines (i.e., d=2) are used to describe Pp over ventilator inspiration, where Tmax represents the total ventilator inspiration time per breath (Tinsp), rounded up to the nearest integer multiple of the defined spline knot width (kw), yielding:

Ppt=j=1M-Ps,jϕj,2t,#5

where Ps,j are the spline coefficients and ϕj,2 are the recursively defined b-spline functions:

ϕj,0(t)=1,Tj<t<Tj+10,otherwise

ϕj,dt=t-TjTj+d-Tjϕj,d-1t+Tj+d+1-tTj+d+1-Tj+1ϕj+1,d-1t, d≥1,#

#6

where Tj are equally spaced points in time with length between points equal to kw=0.05 [s]. Tinsp is defined from where Paw first decreases below PEEP, at the onset of an inspiratory effort, to ventilator cycling. The negative sign in Equation (5), is included to reflect the convention that the patient generates a negative pressure to breath, reducing the pressure at the airways as described by Equation (1).

Aligning with the model formulation first described in [17], the first 75% of the spline coefficients (Ps,j) during an inspiratory effort were constrained to be positive (negative identified Pp) while the latter 25% could be either positive or negative within physiological bounds [17]. The decision to constrain the first 75% of Ps,j aligns with the findings of [26], where it was observed that the strongest performance of the B-Spline model was observed within a range where the first 60-80% of Ps,j were constrained to be positive.

In addition, patient effort is assumed to be absent during expiration. To achieve this, over ventilator expiration, P​p=0.

The model is solved over Nb=20 breaths, yielding constant [E,R] values and unique Pp profiles with each breath. Solving over multiple breaths embeds further domain knowledge into the model, such that changes between breaths are described by Pp, rather than non-physiological changes in respiratory mechanics over short approximately one minute time periods.

2.3.2. B-MIQP Model

In the B-MIQP model, the constraints described in [20] are reformulated to constrain the magnitudes of a set of b-spline functions. The use of b-spline functions to describe Pp reduces the number of decision variables to solve for, such that these constraints can be tractably implemented over multiple breaths. A complete description of the B-MIQP model is given in the supplementary material.

| Table 1: Patient cohort distribution, ventilator settings and B-Spline model outcomes. Results presented as median [inter-quartile range] where appropriate. |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cohort | Pat | PEEP [cmH2O] | PS Above PEEP [cmH2O] | E [cmH2O/L] | R [cmH2O/L/s] | RMSE [cmH2O] | Peak △Pes [cmH2O] | Peak Pp [cmH2O] | PTP p-values |
| 1 | 1a | 6 | 5 | 6.7 | 5.2 | 0.5 | -3.2 [-3.5, -2.9] | -0.7 [-0.8, -0.7] | 3.58E-05 |
|  | 1b | 5 | 6 | 6.8 | 6.1 | 0.8 | -5.8 [-6.0, -5.5] | -0.8 [-1.0, -0.7] | 3.09E-08 |
|  | 2a | 10 | 6 | 10.3 | 3.5 | 0.8 | -5.3 [-5.8, -5.0] | -1.1 [-1.1, -1.0] | 2.65E-08 |
|  | 2b | 12 | 10 | 14.8 | 6.6 | 1.2 | -1.7 [-2.0, -1.4] | -0.9 [-2.3, -0.6] | 4.98E-01 |
|  | 3 | 10 | 10 | 18.1 | 7.2 | 1.5 | -2.3 [-2.5, -2.1] | -1.6 [-2.4, -0.6] | 1.16E-03 |
| 2 | 4 | 7 | 10 | 14.9 | 7.0 | 0.4 | -2.6 [-3.2, -2.4] | -3.2 [-3.6, -3.1] | 2.94E-03 |
|  | 5 | 8 | 8 | 14.9 | 8.8 | 1.7 | -5.4 [-6.1, -4.8] | -3.2 [-3.6, -3.0] | 1.73E-05 |
|  | 6 | 8 | 3 | 5.3 | 1.1 | 0.3 | -5.6 [-5.8, -5.4] | -2.0 [-2.1, -1.7] | 8.44E-05 |
|  | 7 | 12 | 4 | 4.0 | 2.9 | 0.4 | -16.5 [-17.1, -15.2] | -3.1 [-3.7, -2.8] | 1.79E-06 |
|  | 8 | 8 | 10 | 12.3 | 6.6 | 0.5 | -2.2 [-2.4, -1.7] | -1.6 [-1.8, -1.5] | 6.00E-01 |
|  | 9 | 12 | 9 | 7.4 | 4.3 | 0.8 | -9.8 [-10.5, -8.8] | -2.1 [-2.8, -1.8] | 3.85E-07 |
| PEEP: positive end-expiratory pressure; PS: pressure support; RMSE: root mean square error; PTP: pressure time product. |  |  |  |  |  |  |  |  |  |

Briefly, Pp is described by Equation (5), where Ps,j values are constrained so the Pp profile must be monotonically decreasing then increasing, and subsequently equal to 0 over a breath. To this end, successive Ps,j values are constrained to be either monotonically increasing or decreasing; or equal to zero within three defined regions.

A set of switching points (Ts,i) are used to define the boundaries of these regions. The locations of Ts,i are determined within a mixed-integer framework using binary variables and are bounded relative to ventilator triggering and cycling off. As compared to the B-Spline model described in Section 2.3.1, prolonged inspiratory effort can be captured by the model by placement of the expiratory switching point after ventilator cycling. Beyond this point, prior to the onset of the subsequent inspiratory effort, b-spline coefficients (Ps) values are constrained equal to zero.

The revised B-MIQP single compartment lung model is given by:

Pawt=PEEP+EinspVt1-γt+RinspVt1-γt+EexpV(t)γ(t)+RexpV(t)γ(t)+j=1M-Ps,jΦj,2(t),  (7)

where [Einsp, Eexp, Rinsp, Rexp] represent unique sets of inspiratory and expiratory respiratory mechanics. Furthermore, γ represents a scalar binary indicator – equal to either 0 or 1 at each sampled point in time – ensuring inspiratory and expiratory mechanics are considered during ventilator inspiration and expiration respectively [27]. Note that j=1M-Ps,jΦj,2(t) is identically zero during the estimated patient neural expiration.

In this paper, Nb=10 breaths were considered per model fitting (or Nb=9 for Patients 6 and 9) as a compromise between model regularization and solve times. For fair comparison with the B-Spline model, the B-MIQP model was solved twice using two sets of breaths such that both models were fit to all available breaths. Common values for respiratory mechanics (Rinsp,Einsp,Rexp,Eexp) were identified per model fitting, while individual Pp profiles were identified for each breath.

Fittings of the B-Spline and B-MIQP models were obtained considering the minimization of the following cost function:

J=t=tstarttend(Pawt-Pawt)2,#8

where Paw represents the model identified airway pressure signal, described by Equations (1) and (7) respectively. The model and optimization problem were formulated using the YALMIP framework (Taipei, Taiwan) within the MATLAB environment (Mathworks, Natick, MA, USA). The optimization problem was solved using the Gurobi MIQP solver (Houston, TX, USA). Median and Inter Quartile Range (IQR) values for model solve time was 75.9 (45.2, 344) [s] across all patients. Further, sensitivity analysis of model outputs for different Nb and kw has been included in the Supplementary Material.

2.4. Analysis and Validation

The methods were first assessed against established performance metrics [17, 26]. First, the root mean square error (RMSE) of the model residuals was considered as a goodness measure of the model fitting. Second, identified respiratory mechanics values and Pp waveforms were compared to physiological known ranges and trends in literature. Third, mean Pp values were plotted against mean identified alveolar pressure (Pa) values, where Pa(t)=Paw(t)-RinspV(t), to assess if the model outcomes fall along lines of constant trans-pulmonary pressure, as first described by [28].

In this study, the Pes signal was recorded, allowing further comparative validation not considered in prior studies [17, 26]. First, a qualitative comparative assessment of the Pp and △Pes profiles was considered. Second, peak magnitudes of △Pes and Pp were compared to assess model estimation accuracy. Third, Pp and Pes negative areas, representing estimates of PTP, were compared to assess model estimation precision with regards to tracking changes in breathing effort. PTP measures were defined:

PTPinv=t1t2△Pestdt,#9

PTPest=t1t2Ptdt,#10

| Table 2: B-MIQP model outcomes. Results presented as median [inter-quartile range] where appropriate. Two sets of respiratory mechanics and PTP p-values values obtained from two unique Nb = 10 breath model fittings. Elsewhere, results presented as median [inter-quartile range] where appropriate. |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pat | Einsp | Rinsp | Eexp | Rexp | RMSE | Peak △Pes | Peak Pp | PTP p-values |
|  | [cmH2O/L] | [cmH2O/L/s] | [cmH2O/L] | [cmH2O/L/s] | [cmH2O] | [cmH2O] | [cmH2O] |  |
| 1a | [9.2, 5.1] | [8.8, 11.1] | [5.3, 3.6] | [3.5, 2.2] | 0.3 | -3.2 [-3.5, -2.9] | -3.4 [-3.5, -3.2] | [1.12E-01, 2.68E-02] |
| 1b | [9.7, 11.2] | [9.2, 7.4] | [3.0, 3.0] | [1.8, 1.8] | 0.5 | -5.8 [-6.3, -5.5] | -3.6 [-4.1, -3.2] | [1.24E-05, 8.18E-03] |
| 2a | [14.7, 10.5] | [5.3, 6.9] | [7.3, 6.1] | [0.7, 0.3] | 0.4 | -5.3 [-5.8, -5.0] | -3.1 [-3.5, -3.0] | [2.60E-04, 4.59E-02] |
| 2b | [19.6, 8.7] | [10.0, 16.7] | [9.5, 6.8] | [2.0, 0.9] | 0.6 | -1.7 [-2.0, -1.4] | -5.9 [-6.9, -4.1] | [5.91E-01, 6.76E-01] |
| 3 | [23.7, 18.0] | [11.7, 14.2] | [10.0, 10.5] | [0.0, 0.0] | 0.7 | -2.3 [-2.5, -2.1] | -5.3 [-5.9, -5.1] | [5.41E-02, 9.66E-02] |
| 4 | [17.0, 16.8] | [8.1, 9.0] | [3.8, 3.9] | [0.0, 0.0] | 0.2 | -2.6 [-3.2, -2.4] | -3.7 [-4.3, -3.5] | [1.39E-01, 5.00E-03] |
| 5 | [28.5, 28.0] | [22.8, 23.8] | [2.8, 3.1] | [0.0, 0.0] | 0.7 | -5.4 [-6.2, -4.8] | -4.9 [-5.3, -4.3] | [7.42E-04, 4.15E-04] |
| 6 | [10.4, 7.5] | [0.5, 2.5] | [5.0, 3.6] | [0.8, 0.2] | 0.3 | -5.6 [-5.8, -5.4] | -3.3 [-3.6, -3.0] | [2.22E-03, 4.30E-02] |
| 7 | [8.3, 7.4] | [4.3, 6.0] | [1.2, 1.3] | [0.0, 0.0] | 0.4 | -16.5 [-17.1, -15.2] | -5.3 [-5.7, -4.7] | [1.03E-04, 3.08E-05] |
| 8 | [15.0, 14.0] | [7.3, 8.6] | [3.8, 3.4] | [0.6, 0.4] | 0.3 | -2.2 [-2.3, -1.7] | -3.3 [-3.7, -3.1] | [4.79E-01, 7.03E-01] |
| 9 | [14.6, 9.4] | [3.9, 6.2] | [3.0, 3.2] | [0.0, 0.0] | 0.6 | -9.8 [-10.5, -8.8] | -5.5 [-7.2, -4.4] | [1.73E-04, 4.02E-04] |
| RMSE: root mean square error; PTP: pressure time product. |  |  |  |  |  |  |  |  |

where, for each breath t1 and t2 represent the start and end of patient inspiration, respectively.

Considering PTPinv, the start of inspiration was determined as the start of the swing in △Pes prior to ventilator triggering. Comparatively, for B-MIQP identified PTPest, the start of inspiration was defined as the identified inspiratory switching point time (Ts, insp). Further, considering the B-Spline model, t1 was defined as the point where Paw falls below PEEP prior to ventilator triggering. Small differences in t1 between methods have been considered to analyze the extent to which the models can describe the onset of breathing effort. Of note, the use of Ts allows the B-MIQP model greater freedom, as compared to the B-Spline model, to select the starting point of an inspiratory effort. For both measures, the end of inspiration was defined to be the point of ventilator cycling-off. Where △Pes became positive prior to the end of ventilator cycling-off, the end of inspiration was instead defined as the final non-positive value in the △Pes signal.

Linear regression was performed to identify statistically significant relationships between the PTP measures. The y-intercept was not constrained to pass through the origin (0,0) and identified relationships were considered statistically significant where the two-tail p-value was <0.05. Matlab’s fitlm function was used to perform linear regression analysis.

To analyze agreement, Bland-Altman (BA) plots considering model derived (PTPest) and measured (PTPinv) PTP values were considered. To characterize model error, linear trend lines were fit to the BA outcomes, where the y-intercept was unconstrained.

For completeness, we also compare the B-Spline and B-MIQP models to the CO and SO models. Example model fittings for each patient obtained and Pp profiles were compared. As a means to briefly compare models over a single breath window, fittings produced by all models considered in this study were produced, considering a single patient (Patient 2a) and Nb = 1. As recommended in [16], an optimal value for the regularization parameter λ needs to be obtained prior to use of the SO method across a clinical cohort. As such, SO model fits for different values of λ were considered with the aim to identify an appropriate cohort wide value. To complement, as a means to reduce the dimensionality of the optimization problem and evaluate the practical identifiability of the model, B-Spline

Figure 3: Example B-Spline and B-MIQP model fittings for all patients.

model fits for different fixed values of R were also considered [30, 31, 32].

2.4.1 Filtering and shifting the esophageal pressure signal

Prior to analysis, Pes was either filtered or smoothed to remove Cardiogenic Oscillations (CGO) and high frequency noise. As shown in, Figure 1a, CGOs from the heartbeat can distort the Pes signal [33], making it difficult to identify the underlying signal that captures variations in Ppl. Further, the magnitude of CGOs was observed to vary between patients thereby contributing to inter-patient variability in the validation signal. As such, different filtering approaches were considered to appropriately account for these variabilities. Figure 1 provides an example of the two different filtering approaches considered. First, with reference to Figure 1a, a filter adapted from [29] was selectively applied to the Pes signal. Following singular value decomposition, reconstructed components of the unfiltered Pes signal were individually filtered using a 4th order low-pass Butterworth filter (cut-off frequency = 1.4 [Hz]) and a 4th order notch Butterworth filter with a notch frequency equal to the identified fundamental frequency of the CGOs. Considering the fast Fourier transform of the unfiltered Pes signal, the fundamental frequency of the CGOs was identified as the first frequency peak above 0.8 [Hz] which did not overlap with perceived harmonics of the Ppl signal. This filter was applied only where respiratory effort harmonic and CGO fundamental frequencies could be distinguished and CGO magnitude was considered large enough such that inspiratory effort profiles and peaks were significantly distorted. The cut-off value for this classification was 0.5 cmH​2O and was evaluated during neural expiration to visually separate CGOs and respiratory effort. As shown in Figure 1a where the filter has been applied, there is a clear distinction between the Ppl harmonic (2⋅fPpl≈0.5 [Hz]) and the CGO fundamental frequency (fCGO≈1.25 [Hz]). Further, it can be observed that the magnitude of CGOs, as highlighted by the dashed lines,

Figure 4: Mean identified patient effort versus mean identified alveolar pressure. B-Spline Model.

exceed the 0.5 cmH​2O threshold. In comparison, considering Figure 1b, for a number of patients, the observed frequency peaks appeared to be fPpl harmonics rather than fCGO. Further, in some instances, CGO magnitude was <0.5 cmH​2O. In these instances, a moving average filter considering a 20 data point span was instead applied to the unfiltered Pes signal to dampen high frequency noise exclusively. Of note, the filter approaches are applied to the Pes validation signal only and therefore, will not impact the reproducibility of the described models.

3. Results

Examples model fittings, from the B-Spline and B-MIQP models, are given in Figures 2 and 3. Model goodness-of-fit as identified by the Root Mean Squared Error (RMSE) for B-MIQP (0.4 [0.3, 0.6]) and B-Spline (0.8 [0.5, 1.2]) models were similar indicating no loss in model fitting quality with the introduction of additional constraints. Overall, P and △Pes profiles differed, though, the slopes of the signals aligned within sections of the breath. Identified inspiratory respiratory mechanics for both models aligned with physiological expectations (see Table 1 and Table 2). By comparison, B-MIQP model identified expiratory respiratory mechanics were often non-physiological.

Figures 4 and 5 consider plots of mean Pp and mean identified alveolar pressure (Pa). The dashed lines represent constant iso-transpulmonary pressure, iso-Ptp [28]. Considering both models, model outcomes align with iso-Ptp reasonably well, where deviations are likely due to the disconnect between the varying patient-specific demand versus the static inspiratory pressure ventilator support target between breaths [18].

Statistically significant (two-tailed p<0.05) relationships between PTPinv and PTPest were identified for the majority of patients considering both models. P-values are listed in Tables 1 and 2 while identified PTP trends are shown in Figure 6. Overall, a greater proportion of the relationships observed, considering the B-Spline model as compared to the B-MIQP model, could be considered statistically significant. Considering the B-Spline model, the two relationships which were not considered statistically significant (p>0.05) were observed for patients exerting smaller breathing efforts (median peak △Pes≤2.2 [cmH​2O]).

Due to the small number of data points with ∼20 breaths per fitting, individual data points can have a large influence on the identified relationships. Cook’s Distance analysis was performed to identify individual points with the greatest influence on the identified relationships. Any point where its corresponding Cook’s Distance value was greater than 3 times the mean Cook’s distance value were considered outliers if subsequent removal of this point improved the identified correlation between PTPinv and PTPest.

Figure 7 plots R2 and m values obtained from identified PTP trends, against mean △Pes values. R2 values were observed to increase with the magnitude of patient effort. In addition, B-MIQP method identified m values varied to greater degree and tended to be larger than B-Spline method identified m values.

Figure 8 provides example fits considering CO and SO models. Table 3 lists the model outcomes of fittings of all models considering a single breath window (Nb=1).

Figure 5: Mean identified patient effort versus mean identified alveolar pressure. B-MIQP Model.

Considering the CO method, model outcomes, including Pp profiles, peak values and respiratory mechanics were similar to B-MIQP model outcomes. In comparison, SO model fits often varied greatly from other considered models and between model fittings where λ was adjusted. Figure 9, provides an example SO model fits for different λ, and a B-Spline model fits for different R values. It was observed Pp profiles and peak magnitudes varied greatly for varying λ and R values.

Figure 10 and 11 illustrate BA plots of B-Spline and B-MIQP model derived PTPest agreement with PTPinv. Mean difference values highlight the accuracy of the models while the limits of agreement (LOA) assess model precision. Further, fitted linear trends to BA outcomes vary significantly between models and patients in both gradient and correlation. This variation characterizes the relative ability of the models to track changes in SB effort in different clinical scenarios.

4. Discussion

In this study, the performance of the B-Spline and B-MIQP methods for SB effort are analyzed and compared in different clinical scenarios, considering patients receiving PSV. The main findings of the paper were: (1) the B-MIQP model better describes the SB effort profile in comparison to the B-Spline model, as defined by △Pes; (2) both models could often track changes in SB efforts with good precision, but were unable to accurately identify the magnitude of SB effort in all cases; (3) the performance of these methods is dependent on the extent to which the Paw profile is disrupted by the patient effort signal and thus, model performance is likely correlated to level of patient effort and the degree of synchrony between the patient and the ventilator; and (4) model estimation accuracy for all methods considered is likely partially limited by the rapid onset of ventilator support, representing a structural limitation related to PSV.

4.1 Model outputs

Identified model outcomes suggest that both models may yield physiologically reasonable outputs. Considering RMSE listed in Tables 1 and 2, both models produced modeled airway pressure (Paw) fittings with minimal model error. The B-MIQP model generally produced model fits with smaller error, predominantly due to improved fitting to expiration data, achieved using separate inspiratory/expiratory respiratory mechanics.

Identified inspiratory respiratory mechanics for both models (Tables 1 and 2) are generally similar to those obtained in a previous study evaluating the b-spline model on a separate PSV patient cohort [18]. In addition, for a number of patients, the B-MIQP model yielded non-physiological expiratory respiratory mechanics estimates (e.g.,Rexp=0 cmH​2O). For these patients, V did not follow the expected decay pattern and △Pes did not return to 0 [cmH​2O] prior to ventilator cycling. Combined, these trends suggest the presence of extended inspiratory efforts and/or active expiratory effort. Despite this, inspiratory respiratory mechanics identified remained physiologically reasonable. This suggests that by decoupling inspiratory and expiratory respiratory mechanics, extended patient efforts might not deteriorate the quality of the inspiratory respiratory mechanics estimates.

Plots of mean Pp and mean Pa generally fall along dashed lines representing constant mean transpulmonary pressure (Figures 4 and 5). As such, model outcomes reflect expected patient neurological feedback mechanisms. Deviations in outcomes from the dashed lines (namely Patients 3 and 9) likely

Figure 6: PTPinv versus PTPest relationships.

indicate the presence of poor patient-ventilator interactions, as first identified by [18]. Overall, these results provide evidence both methods can yield physiological model outcomes on two separate patient cohorts.

4.2. Ability of models to track changes in patient efforts breath to breath

Measurements of esophageal pressure were obtained in this study, allowing for further validation of the presented models. Comparing Pp and △Pes waveforms in Figure 3, it is clear the overall signal profiles often differ. However, closer inspection reveals the slopes of the signals appear to align within sections of the breath where Paw is disturbed by the patient’s SB effort. The slopes of Pp and △Pes tend to align during the onset of inspiratory effort, and where Paw trends upwards in the latter half of inspiration. Considering the onset of effort, this result is as expected, given the dip in Paw is generated solely by the onset of patient effort. Furthermore, towards the end of the inspiratory phase, the ventilator controller is unable to adjust the flow-rate to perfectly respond to the increase in Ppl (towards 0) as the patient finishes inhalation and as a result, a rapid increase in Paw above the set point is observed. Again, this section of the breath cannot be fully described by ventilator signals and thus must be partially described by Pp.

It was observed that greater efforts tend to result in greater disturbances in Paw. As highlighted in Figure 2, Patient 7 exerts a comparatively larger effort as identified by the peak magnitude of △Pes (see Table 1), resulting in a larger observed rise in Paw towards the end of ventilator inspiration. Overall, these results highlight that both the B-Spline and B-MIQP models most reliably capture patient breathing effort when Paw is visibly disturbed by the patient’s SB. In addition, these results provide initial evidence that both models are able to capture changes in SB effort given the observed relationship between the peak magnitude of effort and the degree to which Paw deviates from the reference PS signal.

Plots of PTPinv and PTPest provide further evidence for each model’s ability to track breath-by-breath changes in SB effort. Linear positive-trending regression curves identified in Figure 6 suggest changes in patient effort as measured by  PTPinv could be captured by changes in PTPest where identified trends could be considered statistically significant (p-value <0.05). Comparison of the coefficient of determination (R2) and the mean peak △Pes (Figure 7), reveals the performance of both models is likely dependent on the level of patient effort. For increasing mean peak magnitudes of △Pes, R2 values trend towards a value of 1. This outcome suggests changes in PTPest better reflect changes in PTPinv where SB efforts are larger in magnitude.

Overall, the results obtained suggest both models may enable clinicians to track changes in breathing efforts on a breath-by-breath basis, where the precision of PTPest to predict variations in PTPinv is likely dependent on the degree to which SB efforts disturb the Paw signal.

Comparison of the two models presented suggests the B-MIQP model may identify a greater amount of information regarding SB effort. First, with reference to Figure 6, the B-MIQP model consistently estimates larger PTP values. Second, as seen in Figure 7, the slopes of the identified regression lines (m) describing the PTP relationships in Figure 6 were generally closer to m=1 indicating the model may capture more information regarding the changes in PTP between breaths. Third, for multiple patients, it can be observed that B-MIQP Pp signals follow slopes of the △Pes signal after the point where Paw reaches the PS target. In comparison, the B-Spline Pp signal slopes align with △Pes during this period only where the SB effort generates a disturbance in the Paw signal. However, a comparatively smaller number of B-MIQP model identified PTP relationships could be considered statistically significant as compared to the B-Spline model (see p-values in Tables 1 and 2). Variations in PTPinv that could not be described by the B-MIQP model derived PTPest likely did not align with the assumed breathing profile. Therefore, these results suggest implementation of further Pp shape assumptions may improve model convergence towards a physiologically interpretable solution. However, this may occur at the expense of reduced model flexibility, such that effort profiles that deviate from the assumed model structure are not fully captured.

In terms of model precision, both models appear to track changes in SB effort with similar uncertainty. As seen in Figure 6 and 10, for larger SB efforts, the B-MIQP model tends to underestimate PTPinv, while for smaller efforts, PTPinv is overestimated. In comparison, the B-Spline model consistently underestimates PTPinv (see Figure 11). With reference to Figure 7, absolute m values vary within a tighter range comparing the B-Spline (0.1 [0.1, 0.2]) and B-MIQP (0.6 [0.3, 0.8]) models. However, consideration of the percentage difference of the IQR with respect to the median m value ([-12.3%, 62.9%] and [-43.4%, 51.5%] for the B-Spline and B-MIQP models respectively) suggests neither model appears to track changes in effort with greater precision.

Figure 7: Mean peak esophageal pressure versus coefficient correlation (top), and boxplot of slopes of identified linear regression trends from Figure 6 (bottom).

Consideration of model fittings where Nb=1 revealed that the spline based models approach similar solutions compared to fittings where Nb > 1 (model outcomes listed in Table 3). Further, the B-MIQP model solution was observed to align with the CO and SO model outcomes. This was expected given the similarity in the models’ constraints (i.e., a piecewise monotonic SB constraint inherently implies a degree of smoothness in the SB effort profile). As such, this highlights the strength of the spline based models in their ability to estimate SB effort over multiple breaths, reducing the variation in estimated respiratory mechanics (see Table 4 in the Supplementary Material), thereby improving model precision. Overall, the findings in this paper suggest both models may be able to track changes in patient effort with similar precision, where performance is likely improved for greater SB efforts and where assumed non-physiological variation in identified respiratory mechanics, over short periods of time, is reduced.

BA plots in Figures 10 and 11 further characterize model error. Fitted trend lines with a gradient (m) or R2 values near zero highlight patients where the models were more likely to track changes in PTPinv (△PTPinv) with similar precision over varying magnitudes of SB effort. Comparatively, for model fittings where m > 0, a component of △PTPinv may not be captured by PTPest (see for example Patient 2a in both figures). Similarly, considering the B-MIQP model, for patients where Pp overestimates △Pes, it was observed m < 0, suggesting △PTPinv was overestimated for increasing SB effort (see Patient 3 in Figure 10). Comparing models, where SB efforts were larger, B-MIQP related R2 values were smaller, suggesting the model’s greater ability to track △PTPinv. This result further reflects the finding that the B-MIQP model may be able to capture a greater proportion of inspiratory effort. Conversely, for smaller efforts, B-Spline R2 values were smaller and m was closer to zero suggesting better tracking. As such, these results highlight the potential greater precision of the B-Spline method, due to its more relaxed constraints, where SB efforts did not align with model assumptions.

These results demonstrate that clinicians may be able to monitor SB effort over time utilizing these models in a clinical setting, where model reliability is governed by the magnitude of SB effort and the extent to which effort profiles align with model assumptions.

Figure 8: Example fittings to a singular breath considering Constrained Optimization (top) and Sparse Optimization (bottom) models. λ = 2.5 × 10−3 considered, as prescribed by [16].

4.3. Accuracy of breathing effort estimates

| Table 3: Single breath window model fitting outcomes. |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Peak effort PTP Elastance Resistance [cmH2O] [cmH2O·s] [cmH2O/L] [cmH2O/L/s] |  |  |  |
| △Pes | -5.3 (-5.8, -5.3) | 2.1 (1.9, 2.5) | - | - |
| B-Spline | -1.4 (-1.4, -1.5) | 0.4 (0.4, 0.5) | 11 (11, 12) | 4.9 (4.8, 5.0) |
| B-MIQP | -6.1 (-7.0, -5.6) | 3.3 (3.0, 3.7) | 24 (21, 25) | 6.1 (5.8, 6.6) |
| CO | -3.4 (-3.8, -3.2) | 1.8 (1.7, 2.0) | 14 (14, 15) | 6.9 (6.7, 7.2) |
| SO | -5.9 (-5.3, -6.9) | 3.3 (3.0, 3.8) | 21 (21, 24) | 6.7 (6.6, 6.8) |

It was observed that the B-spline model could estimate peak Ppl accurately only where peak efforts occurred at or close to ventilator triggering. See for example, Patient 8, as reflected by model tracings in Figure 3 and mean difference values in Figure 11. In comparison, accuracy of B-MIQP model estimates was observed to be influenced by the magnitude of patient effort, as well as the synchrony between the patient and ventilator. Visually, considering Figure 3, the best agreement between Ppl and Ppl profiles was observed where efforts were larger and the swing in Pes finished prior to ventilator cycling (see Patients 1a, 2a, 5, and 6). Patient 2b exerts small efforts and as a consequence, Ppl tended to overestimate Ppl. It is likely that in these instances, there is insufficient variation in the Paw signal to accurately identify the underlying Ppl signal. Alternatively, it is possible that the model instead provided an estimate for Pmus, given that (3) is also a valid solution to (1). In this instance, △Pes is observed to trend upwards above 0 [cmH​2O] at end-inspiration. If Ppl becomes positive following a similar trend, the Ppl physiological solution described by Equation (2) is not contained within the feasible set. As described in Section 2.3.2, B-MIQP model derived P signals are constrained to be non-positive and therefore cannot capture positive components of the breathing effort signal. As such, the B-MIQP model for similar breathing efforts was instead more likely to converge to the Pmus physiological solution described by Equation (3). Further, the downwards trends observed in Figure 10 could signify instances where P​p estimates Pmus and, therefore, the variation in Pmus between breaths exceeds variation in Ppl.

The accuracy of the estimates is also impacted where – typically larger – efforts extend beyond ventilator cycling, as is the case for Patient 7 and, to a lesser extent, Patient 9. As initially discussed in [19], respiratory mechanics are uniquely identifiable over the latter portion of the breath in which SB effort ceases prior to ventilator cycling. Overall, these outcomes suggest the accuracy of the B-MIQP model may be dependent on the degree to which Paw is disturbed by Ppl, characterized by the magnitude of efforts and the extent to which SB efforts ceases  prior to ventilator cycling. As such, where significant Paw variation due to SB effort is present, clinicians may be able to utilize the B-MIQP model to non-invasively and continuously estimate SB effort.

A shared model limitation has been identified that appears to partially limit the accuracy of model estimations. For many patients considering both models, P was observed to plateau after ventilator triggering (Figure 3). This trend does not align with the downwards slope of △Pes during this period nor does it meet physiological expectations.

Referencing Figure 8, highlighting Patient 6, it can be seen that the same model limitation may also be observed in the CO and SO models. For all model fittings produced by the CO model, a plateau during the onset of ventilator support was observed (see Supplementary Material). Similarly, the slopes of Pp signals generated by the SO model did not align with the slopes of △Pes during the onset of patient effort and the ramp up of ventilator support. This trend remained consistent for all λ values considered (see Supplementary Material xxx).

This observed trend aligns with the results from [32]. In their study, the authors identified patient effort pressure profiles from ventilator signal data by assuming the respiratory mechanics. Where the assumed resistance value was less than the true value, a plateau in the identified effort profile was observed. Given the plateaus observed, it is likely both models underestimate Rinsp and, given their interrelated nature, the peak magnitude of patient effort is also underestimated.

It is hypothesized the likely underestimation of resistance and SB effort is due to the rapid onset of ventilator support, effectively concealing information about the patient effort pressure profile. At the point of triggering, the ventilator controller rapidly increases the flow rate such to track the rise in target airway pressure. Considering the B-Spline model fits, this rapid rise in Paw can be predominately described by the resistance-flow pressure (Pflow=RV). As seen in Figure 2, for both patients, Pp rapidly returns to 0 [cmH​2O], appearing to plateau until Paw begins to rise above the target PS. Therefore,

(a) B-Spline Model

(b) Sparse Optimization Model

Figure 9: B-Spline model fitting for different R values (left) and Sparse Optimization fitting for different λ values (right).

over this period, Paw can be described solely by the V and V signals.

As a means of reducing the complexity of the problem presented in this paper, to overcome the observed limitation, multiple B-Spline model fittings were performed for different fixed R values, similar to the work of [30]. As shown in Figure 9, for increasing values of R, the observed plateau in Pp became less prominent. In addition, model fits with respect to Paw did not appear to visually worsen for increasing R. However, for increasing R, the magnitude of noise appears to increase, worsening the optimization cost value.

Comparing to the SO method, considering Figure 9, it can be seen that both the peak magnitude and profile of Pp change significantly for different values of λ. Further, it appears that no single λ value exists that yields accurate estimates for all patients (see Supplementary Material: Appendix A).

Overall, for all methods considered, it can be concluded that for SB efforts that yield limited variation in Paw, prior knowledge of R and/or the shape of Pp during the ventilator support ramping phase may be required to improve model identification. Likewise, for patients characterized by comparatively elevated resistance values (i.e., ARDS patients), model estimation accuracy may be disproportionately impacted by measurement noise in the V signal.

There appears to be no universally superior model-based method. Each model represents a unique balance between model flexibility and convergence towards physiologically interpretable solutions. The CO method makes similar assumptions to the B-MIQP method and, therefore, is likely to converge to similar physiological solutions and will similarly be unable to capture variations in breathing effort that do not align with the model constraints. The SO method is comparatively less constrained, where model flexibility is scalable with λ. Notably, as λ is increased, the SO model is likely to be increasingly sensitive to noise, given the consideration of the patient breathing profile’s second derivative. Further, for all methods considered, a structural limitation related to PSV partially limited the accuracy of the model outputs. Ventilator waveform profiles vary significantly between ventilation modes and as such, the performance of these methods is likely to – and previously has been shown to – differ across non-invasive [34], controlled [16], and assist [17, 18] modes. It is likely then that the most appropriate method for any given clinical scenario will be dependent on the state of the patient and their synchrony with the ventilator, the degree of noise in the measured ventilator signals, and importantly the ventilation mode.

Figure 10: Per patient Bland-Altman plots considering pressure time product (PTP) estimates obtained from B-MIQP model.

Figure 10: Per patient Bland-Altman plots considering pressure time product (PTP) estimates obtained from B-MIQP model.

4.4. Limitations

The B-MIQP and B-Spline models were validated against esophageal pressure derived measures. These measures have been used to validate other methods to estimate patient effort [7, 8, 12, 14]. Despite this, the validity of Pes as a surrogate measure of Ppl has historically been somewhat contested and as such, a degree of uncertainty exists [35, 36]. However, more recent studies have addressed methods to correctly establish and analyze the measures such to account for these limitations [37]. Despite its potential shortcomings, esophageal pressure measurements remain the gold standard [33].

Given the careful selection of breaths to avoid Pes inaccuracies, only a limited number of breaths per patient were available for analysis. As such, the impacts of potential outliers on identified trends were exacerbated. Despite this, statistically significant relationships could still often be identified, especially where Paw was significantly disturbed by breathing efforts.

For both spline-based models active expiratory effort was assumed absent. It is envisioned that a clinical implementation of the spline-based models would include a preliminary asynchrony detection algorithm to identify and remove asynchronous efforts prior to model fitting, thereby mitigating potential model limitations related to violations of the assumed breathing profile and reduced variation in Paw. Asynchronies can be identified by deviations from expected ventilator waveform profiles, as described in [38, 39].

Improved performance of the B-MIQP model was consistently observed where breathing efforts ceased prior to ventilator cycling. Due to a lack of available relevant data, a more systematic analysis of the method, considering synchronous breaths only, could not be performed. Further, Ecw estimates could not be derived from Pes measurements for all patients, preventing comparison of Pp to Pmus estimates. As identified in Section 4.3, B-MIQP model outputs may at times better reflect the muscle pressure solution to the model. Future work will consider further model development to ensure convergence towards a singular physiological solution. In addition, the sample size considered in this paper was limited due to the scarcity of patient data where Pes measurements were recorded. Despite this, the performance of the models in different clinical scenarios could still be effectively evaluated. Additional future work will consider a larger systematic analysis of the B-MIQP model limited to breaths synchronous with the ventilator and sufficient in magnitude to generate significant variation in Paw. Further, this analysis will consider

Figure 11: Per patient Bland-Altman plots considering pressure time product (PTP) estimates obtained from B-Spline model.

model evaluation against both Pes-derived Ppl and Pmus measures.

5. Conclusions

In this paper, we show that both the B-Spline and B-MIQP models may be able to provide a non-invasive measure to track changes in spontaneous breathing effort. In a clinical setting, these models may be able to aid clinicians in adjusting ventilator settings and/or sedative dosage protocols to maintain SB effort that promotes LDPV. Validation of the models was performed considering esophageal pressure measurements; the current gold standard measure. Model reliability was observed to be linked to the extent to which the airway pressure signal was disturbed by the patient’s breathing efforts. A shared model limitation across all considered methods was identified that reduced estimation accuracy. To overcome this limitation, a greater degree of domain knowledge is likely required to ensure complete model identifiability. Overall, this study provides clear characterization of the potential use-cases for model-based methods, whilst acknowledging identified model limitations that reveal a clear pathway for further development.

References

- M. Bertoni, S. Spadaro, E. Goligher, Monitoring patient respiratory effort during mechanical ventilation: lung and diaphragm-protective ventilation, Annual Update in In-tensive Care and Emergency Medicine 2020 (2020) 21–35.

- T. Pettenuzzo, N. Sella, F. Zarantonello, A. D. Cas-sai, F. Geraldini, P. Persona, E. Pistollato, A. Bos-colo, P. Navalesi, How to recognize patients at risk of self-inflicted lung injury, Expert review of respiratory medicine 16 (9) (2022) 963–971.

- E. C. Goligher, M. Dres, E. Fan, G. D. Rubenfeld, D. C. Scales, M. S. Herridge, S. Vorona, M. C. Sklar, N. Rit-tayamai, A. Lanys, et al., Mechanical ventilation–induced diaphragm atrophy strongly impacts clinical outcomes, American journal of respiratory and critical care medicine 197 (2) (2018) 204–213.

- T. Schepens, M. Dres, L. Heunks, E. Goligher, Diaphragm-protective mechanical ventilation, Current Opinion in Critical Care 25 (1) (2019) 77–85.

- E. C. Goligher, M. Dres, B. K. Patel, S. K. Sahetya, J. R. Beitler, I. Telias, T. Yoshida, K. Vaporidi, D. L. Grieco,

T. Schepens, et al., Lung-and diaphragm-protective ventilation, American journal of respiratory and critical care medicine 202 (7) (2020) 950–961.

- T. Mauri, T. Yoshida, G. Bellani, E. C. Goligher, G. Carteaux, N. Rittayamai, F. Mojoli, D. Chiumello, L. Piquilloud, S. Grasso, et al., Esophageal and transpulmonary pressure in the clinical setting: meaning, usefulness and perspectives, Intensive care medicine 42 (2016) 1360–1373.

- F. Albani, L. Pisani, G. Ciabatti, F. Fusina, B. Buizza, A. Granato, V. Lippolis, E. Aniballi, F. Murgolo, A. Rosano, et al., Flow index: a novel, non-invasive, continuous, quantitative method to evaluate patient inspiratory effort during pressure support ventilation, Critical Care 25 (1) (2021) 1–11.

- M. Bertoni, I. Telias, M. Urner, M. Long, L. Del Sorbo,

E. Fan, C. Sinderby, J. Beck, L. Liu, H. Qiu, et al., A novel non-invasive method to detect excessively high respiratory effort and dynamic transpulmonary driving pressure during mechanical ventilation, Critical Care 23 (2019) 1–10.

- I. Telias, F. Damiani, L. Brochard, The airway occlusion pressure (p0.1) to monitor respiratory drive during mechanical ventilation: increasing awareness of a not-so-new problem, Intensive care medicine 44 (2018) 1532–1535.

- T. Yoshida, M. B. Amato, D. L. Grieco, L. Chen, C. A. Lima, R. Roldan, C. C. Morais, S. Gomes, E. L. Costa, P. F. Cardoso, et al., Esophageal manometry and regional transpulmonary pressure in lung injury, American journal of respiratory and critical care medicine 197 (8) (2018) 1018–1026.

- A. H. Jonkman, H. J. De Vries, L. M. Heunks, Physiology of the respiratory drive in icu patients: implications for diagnosis and treatment, Critical Care 24 (1) (2020) 104.

- I. Telias, D. Junhasavasdikul, N. Rittayamai, L. Piquil-loud, L. Chen, N. D. Ferguson, E. C. Goligher, L. Brochard, Airway occlusion pressure as an estimate of respiratory drive and inspiratory effort during assisted ventilation, American journal of respiratory and critical care medicine 201 (9) (2020) 1086–1098.

- R. Gao, J.-X. Zhou, Y.-L. Yang, S.-S. Xu, Y.-M. Zhou,

L. Zhang, M.-Y. Miao, Use of pressure muscle index to predict the contribution of patient’s inspiratory effort during pressure support ventilation: a prospective physiolog-ical study, Frontiers in Medicine 11 (2024) 1390878.

- M. Docci, E. Rezoagli, M. Teggia-Droghi, A. Coppadoro,

M. Pozzi, A. Grassi, I. Bianchi, G. Foti, G. Bellani, In-dividual response in patient’s effort and driving pressure to variations in assistance during pressure support ventilation, Annals of Intensive Care 13 (1) (2023) 132.

- M. Miao, W. Chen, Y. Zhou, R. Gao, D. Song, S. Wang,

Y. Yang, L. Zhang, J. Zhou, Validation of the flow index to detect low inspiratory effort during pressure support ventilation, Annals of Intensive Care 12 (1) (2022) 89.

- J. Reinders, B. Hunnekens, N. van de Wouw, T. Oomen, Noninvasive breathing effort estimation of mechanically ventilated patients using sparse optimization, IEEE Open Journal of Control Systems 1 (2022) 57–68.

- J. L. Knopp, E. Guy, K. T. Kim, G. M. Shaw, J. G. Chase, B-spline modelling of inspiratory drive in nava-ventilated patients, IFAC-PapersOnLine 54 (15) (2021) 103–108.

- Q. Sun, J. G. Chase, C. Zhou, M. H. Tawhai, J. L. Knopp,

K. Möller, G. M. Shaw, T. Desaive, Estimating patient spontaneous breathing effort in mechanical ventilation using a b-splines function approach, IFAC Journal of Systems and Control (2024) 100259.

- F. Vicario, A. Albanese, N. Karamolegkos, D. Wang,

A. Seiver, N. W. Chbat, Noninvasive estimation of respiratory mechanics in spontaneously breathing ventilated patients: a constrained optimization approach, IEEE Trans-actions on Biomedical Engineering 63 (4) (2015) 775–

787.

- M. H. Victor Jr, M. R. Maximo, M. M. Matsumoto, S. M. Pereira, M. R. Tucci, Mixed-integer quadratic program-ming approach for noninvasive estimation of respiratory effort profile during pressure support ventilation, International Journal for Numerical Methods in Biomedical Engineering 39 (1) (2023) e3668.

- J. H. Bates, Lung mechanics: an inverse modeling approach, Cambridge University Press, 2009.

- J. Knopp, G. Chase, K. Kim, G. Shaw, Model-based estimation of negative inspiratory driving pressure in patients receiving invasive nava mechanical ventilation, Com-puter Methods and Programs in Biomedicine 208 (2021) 106300.

- K. Lindup, J. Chase, C. Zhou, M. Bertoni, F. Padula,

A. Visioli, Non-invasive patient breathing effort identification: a b-spline and mixed integer solution, IFAC-PapersOnLine 58 (24) (2024) 187–192.

- J. L. Knopp, Y. S. Chiew, D. Georgopoulos, G. M. Shaw,

J. G. Chase, Ubiquity of models describing inspiratory effort dynamics in patients on pressure support ventilation, IFAC Journal of Systems and Control 27 (2024) 100250.

- V. J. Major, Y. S. Chiew, G. M. Shaw, J. G. Chase, Biomedical engineer’s guide to the clinical aspects of in-tensive care mechanical ventilation, Biomedical engineering online 17 (1) (2018) 1–31.

- Q. Sun, J. G. Chase, C. Zhou, M. H. Tawhai, J. L. Knopp,

K. Möller, T. Hu, G. M. Shaw, Patient spontaneous effort estimation in digital twin model with b-spline function, IFAC-PapersOnLine 56 (2) (2023) 2096–2101.

- S. L. Howe, J. G. Chase, D. P. Redmond, S. E. Mor-ton, K. T. Kim, C. Pretty, G. M. Shaw, M. H. Tawhai,

T. Desaive, Inspiratory respiratory mechanics estimation by using expiratory data for reverse-triggered breathing cycles, Computer Methods and Programs in Biomedicine 186 (2020) 105184.

- C. Sinderby, J. Beck, J. Spahija, M. De Marchie,

J. Lacroix, P. Navalesi, A. S. Slutsky, Inspiratory muscle unloading by neurally adjusted ventilatory assist during maximal inspiratory efforts in healthy subjects, Chest 131 (3) (2007) 711–717.

- S. K. Mukhopadhyay, M. Zara, I. Telias, L. Chen,

R. Coudroy, T. Yoshida, L. Brochard, S. Krishnan, A singular spectrum analysis-based data-driven technique for the removal of cardiogenic oscillations in esophageal pressure signals, IEEE Journal of Translational Engineering in Health and Medicine 8 (2020) 1–11.

- E. F. Guy, J. G. Chase, J. L. Knopp, G. M. Shaw, Quanti-fying ventilator unloading in cpap ventilation, Computers in Biology and Medicine 142 (2022) 105225.

- C. Schranz, P. D. Docherty, Y. S. Chiew, J. G. Chase,

K. Möller, Structural identifiability and practical applicability of an alveolar recruitment model for ards patients, IEEE Transactions on biomedical engineering 59 (12) (2012) 3396–3404.

- M. Younes, L. Brochard, S. Grasso, J. Kun, J. Mancebo,

M. Ranieri, J.-C. Richard, H. Younes, A method for monitoring and improving patient: ventilator interaction, In-tensive care medicine 33 (2007) 1337–1346.

- A. H. Jonkman, I. Telias, E. Spinelli, E. Akoumianaki,

L. Piquilloud, The oesophageal balloon for respiratory monitoring in ventilated patients: updated clinical re-view and practical aspects, European Respiratory Review 32 (168) (2023).

- F. Vicario, A. Albanese, D. Wang, N. Karamolegkos,

N. W. Chbat, Constrained optimization for noninvasive estimation of work of breathing, in: 2015 37th Annual Inter-national Conference of the IEEE Engineering in Medicine and Biology Society (EMBC), IEEE, 2015, pp. 5327–5330.

- J. G. Chase, K. Moeller, G. M. Shaw, C. Schranz, Y. S. Chiew, T. Desaive, When the value of gold is zero, BMC research notes 7 (2014) 1–3.

- G. Hedenstierna, Esophageal pressure: benefit and limita-tions., Minerva anestesiologica 78 (8) (2012) 959–966.

- T. Pham, I. Telias, J. R. Beitler, Esophageal manometry, Respiratory Care 65 (6) (2020) 772–792.

- F. Mojoli, M. Pozzi, A. Orlando, I. M. Bianchi, E. Arisi,

G. A. Iotti, A. Braschi, L. Brochard, Timing of inspiratory muscle activity detected from airway pressure and flow during pressure support ventilation: the waveform method, Critical Care 26 (1) (2022) 32.

- L. van de Kamp, J. Reinders, B. Hunnekens, T. Oomen,

N. van de Wouw, Automatic patient-ventilator asynchrony detection framework using objective asynchrony definitions, IFAC Journal of Systems and Control 27 (2024) 100236.